## Standardization Record
C173 Air Meter Standardization
Name of Technician: |  |  | John Piper |  |  | Air Meter ID: |  | AM-1
[6 empty] | Calibrated Cup ID: |  | CC-2
Date of Standardization: |  |  | 2013-08-02
[6 empty] | Balance ID (readable to 0.1 g for bowl): |  | 1177847
[6 empty] | Balance ID (readable to 0.01 g for cup and neck): |  | O-48
Procedure: |  |  | IHC-13 |  |  | Thermometer ID (readable to 0.1 C/F): |  | TMP-2
Bowl Volume Standardization | [5 empty] | Calibrated Cup Volume Standardization
Mass of the Bowl and Plate (g): |  |  |  | 2617.2 |  | Mass of the Calibrated Cup and Plate (g): |  |  | 278.05
Mass of the Bowl, Plate, and Water (g): |  |  |  | 4735.5 |  | Mass of the Calibrated Cup, Plate, and Water (g): |  |  | 299.91
Mass of Water (g): |  |  |  | 2118.3 |  | Mass of Water (g): |  |  | 21.86
Temperature of Water (C or F): |  |  |  | 72 |  | Temperature of Water (C or F): |  |  | 71.4
Density of Water in lb/ft3: |  |  |  | 62.2851 |  | Density of Water in lb/ft3: |  |  | 62.2899
Volume of the Bowl (ft3): |  |  |  | 0.075 |  | Volume of the Cup (ft3): |  |  | 0.00077
Minimum Bowl Volume | [5 empty] | Calibrated Cup Conformance
[6 empty] | Percent Volume of Calibrated Cup: |  |  | 1.03
Is the volume of the bowl at least 0.075 ft3? |  |  |  | Yes |  | Is the Calibrated Cup 0.96 - 1.04%? |  |  | Yes
Check of Graduations in the Neck
|  |  | Temperature of the Water Being Added: |  |  | 72.3
|  |  | Density of the Water Being Added: |  |  | 62.283
|  |  | The Mass of Water That Needs To Be Used for Exactly 1.0% Additions (g): |  |  | 21.18
Nominal Graduation | Estimated Water Level Reading | Estimation Within 0.1% of Graduation? | Estimation Within 0.1% of Cumulative % Added? | Mass of Water Added for Checking the Next Graduation (g) | Percent Added By Volume Based on the Mass | Is the % Added 1.0 Percent of the Bowl? | Cumulative Mass Added | Cumulative % Added | Cumulative % Added Within 0.1 % of Next Graduation?
9 | 9 | Yes | Yes | 21.86 | 1.03 | Yes | 21.86 | 1.03 | Yes
8 | 8 | Yes | Yes | 21.01 | 0.99 | Yes | 42.87 | 2.02 | Yes
7 | 7 | Yes | Yes | 21.54 | 1.02 | Yes | 64.41 | 3.04 | Yes
6 | 6 | Yes | Yes | 20.98 | 0.99 | Yes | 85.39 | 4.03 | Yes
5 | 5 | Yes | Yes | 20.77 | 0.98 | Yes | 106.16 | 5.01 | Yes
4 | 4 | Yes | Yes | 21.81 | 1.03 | Yes | 127.97 | 6.04 | Yes
3 | 3 | Yes | Yes | 20.97 | 0.99 | Yes | 148.94 | 7.03 | Yes
2 | 2 | Yes | Yes | 21.78 | 1.03 | Yes | 170.72 | 8.06 | Yes
1 | 1 | Yes | Yes | 21.82 | 1.03 | Yes | 192.54 | 9.09 | Yes
0 | 0 | Yes | Yes | ---- | ---- | ---- | ---- | ---- | ----
[8 empty] | Form Preparation/Revision Date: | 2024-04-11
Note on I17: Pete Holter: Temperature should be 21.1C or 70F per Section 5.4 of C173-16
Note on F30: Pete Holter: For this part of the standardization, a pitcher of water should be preprared so that this water can be poured into a smaller vessel that will be used to add the water to the meter in the required increments. The calibrated cup can be used as this smaller vessel if it has sufficient volume. But if, for example, the calibrated cup has a volume of only 0.96 percent of the volume of the bowl, its volume is too small to be used for this part of the standardization and a separate vessel will be needed. Determine the temperature of the water in the pitcher so that the calculated percentages being added can be based on the volume. The mass of the water being added each time should be within +/- 0.2 grams of the target indicated in Cell I32.
Note on D35: Pete Holter: Entering a result here is confirmation that you have added water to bring the water level to within +/- 0.1 percent of 9.0 percent as your starting point.
Note on F35: Pete Holter: Unlike the cells below, this cell simply compares the starting water level with the nominal graduation.
Note on G35: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on D36: Pete Holter: This is the estimated water level observed after adding the increment noted in the row above in cell G32. And so on.
Note on F36: Pete Holter: This cell and those below it are comparing the Estimated Water Level in the current row with the Cumulaitve % Added from the previous row.
Note on G36: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on D37: Pete Holter: This is the estimated water level observed after adding the increment noted in the row above in cell G33. And so on.
Note on G37: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G38: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G39: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G40: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G41: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G42: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.
Note on G43: Pete Holter: This is the mass of water being added to go from 9.0 percent to 8.0 percent. The observed reading after the addition of this increment is recorded in the next row in cell E33. And so on.

## Water Density Table
TABLE 3 Density of Water
Temperature |  | kg/m3 | lb/ft3
C | F
15.6 | 60 | 999.01 | 62.366
18.3 | 65 | 998.54 | 62.336
21.1 | 70 | 997.97 | 62.301
23 | 73.4 | 997.54 | 62.274
23.9 | 75 | 997.32 | 62.261
26.7 | 80 | 996.59 | 62.216
29.4 | 85 | 995.83 | 62.166