## Lab Data
D7928-21
Standard Test Method for Particle-Size Distribution (Gradation) of Fine-Grained Soils Using the Sedimentation (Hydrometer) Analysis
Project Information and Field Data
Project Identification: | [8 empty] | Boring Number:
Name of Field Technician: | [8 empty] | Sample Number:
Date of Sampling: | [8 empty] | Depth:
Preliminary Laboratory Data
Date Sample Was Received:
Test Number: | [8 empty] | Test Start Date:
Name of Technician: | [8 empty] | Test End Date:
Prior Testing Performed on the Specimen:
Description and Classification of Soil:
Description of Any Material Excluded:
Any Problems Encountered:
Equipment IDs
Thermometric Device: | [8 empty] | Sedimentation Cylinder: |  |  | Cyl-1
Thermometric Device Resolution: |  |  | 0.1 | [5 empty] | Control Cylinder (if used): |  |  | n/a
Balance readable to 0.01 g: | [8 empty] | Hydrometer: |  |  | Hydro-1
Wet Washing Sieve: | [8 empty] | Hydrometer Type: |  |  | 152H
Specimen Mixing Container: | [8 empty] | Time Measuring Device:
Water Content Specimen Container: | [8 empty] | Oven:
Soil-Suspension Oven Drying Container: | [8 empty] | Other Balance (if needed):
Plus No. 200 Material Drying Container: | [8 empty] | Other Balance Resolution (g): |  |  | 5
Sample Preparation
Sample
99% of the Sample Passes the Following Sieve: |  |  |  |  | 3⁄4 in.
Mass of Sample: |  |  |  |  | 4001 |  | g
Reduced Sample
The Reduced Sample Passes the Following Sieve: |  |  |  |  | 3⁄8 in.
Mass of Reduced Sample: |  |  |  |  | 100 |  | g | ← Sample Mass is too small. Refer to Table 1 worksheet for minimum sample mass requirements.
Sedimentation Sample
Sedimentation Sample Preparation Method: |  |  |  |  | Air-Dried
Separation Sieve: |  |  |  |  | No. 10 (2.0 mm)
Mass of Sedimentation Sample: |  |  |  |  | 45 |  | g | ← Sample Mass is too small. Refer to Table 1 worksheet for minimum sample mass requirements.
Estimated Maximum Moist Mass | [8 empty] | Water Content Determination
Hydrometer Capacity: |  |  |  |  | 55 | g |  |  | Is there enough material for a water content specimen? | [6 empty] | 2
Estimated Percentage of Material Passing the No. 200 (75-µm) Sieve, Nearest 1 %: |  |  |  |  | 45 | % |  |  | Mass of Moist Mass Specimen Container, 0.01 g: |  |  |  |  | 116.45 | ← Not Used
[9 empty] | Mass of Dry Mass Specimen Container, 0.01 g: |  |  |  |  | 116.45 | g
Estimated Water Content, Nearest 1 %: |  |  |  |  | 20 | % |  |  | Mass of Container and Moist Specimen, 0.01 g: |  |  |  |  | 180.54 | ← Using Cell I47
Estimated Moist Mass (Mmest), 1 g: |  |  |  |  | 147 | g |  |  | Mass of Container and Oven Dry Specimen, 0.01 g: |  |  |  |  | 167.04 | g ← Includes MDisp
[9 empty] | Water Content (Wc), Nearest 0.1 %: |  |  |  |  | 49.7 | %
Specific Gravity of the Soil Particles
|  | Soil Specific Gravity (Gs): |  |  | 2.557 |  |  |  |  | Method of Determination: |  |  | Measured
Sedimentation (Hydrometer) Specimen and Sodium Hexametaphosphate
Moist Mass of Sedimentation Specimen (Mm), Nearest 0.01 g : | [5 empty] | 68.27 | g |  | Mass of Dispersant (MDisp), Nearest 0.01 g: |  |  |  |  | 5 | g
Temperature-Density Correction
How are you determining the Temperature-Density Correction? | [11 empty] | 2
Hydrometer Test Data
Elapsed Minutes | Hydrometer Reading | Temperature (°C or °F) |  | Temperature Density Correction Offset | [6 empty] | Effective Depth (cm) |  | Diameter (mm) | Mass Percent Finer
|  |  |  | Companion Measurements |  |  |  | Calibration Relationship
|  |  |  | Hydrometer |  | Temp (°C or °F)
1 | 40 | 20 |  | 5 |  | 20 |  | 6.2 |  |  | 8.4 |  | 0.0406 | 75.8
2 | 36.25 | 20 |  | 5 |  | 20 |  | 6.2 |  |  | 9 |  | 0.0297 | 67.4
4 | 32 | 21.2 |  | 4 |  | 21.1 |  | 5.8 |  |  | 9.7 |  | 0.0218 | 58.8
8 | 29 | 20 |  | 4 |  | 21.1 |  | 6.2 |  |  | 10.2 |  | 0.0158 | 51.1
15 | 25.5 | 21.2 |  | 4 |  | 21 |  | 5.8 |  |  | 10.8 |  | 0.0119 | 44.2
30 | 19.75 | 21.2 |  | 4 |  | 21 |  | 5.8 |  |  | 11.7 |  | 0.0088 | 31.3
60 | 17 | 21.2 |  | 4.25 |  | 21 |  | 5.8 |  |  | 12.2 |  | 0.0063 | 25.1
240 | 15 | 21.1 |  | 4.25 |  | 21.4 |  | 5.8 |  |  | 12.5 |  | 0.0032 | 20.6
1440 | 8.75 | 21.1 |  | 4 |  | 21.2 |  | 5.8 |  |  | 13.5 |  | 0.0014 | 6.6
Wash Data and % Passing the No. 200 Sieve
[7 empty] | Mdd: | 50.59 | g
[7 empty] | Md: | 45.59 | g
[7 empty] | Mdr: | 21.65 | g
[7 empty] | Pp: | 52.5 | %
|  | 0.075 | 52.5
|  | 0.0406 | 75.8
|  | 0.0297 | 67.4
|  | 0.0218 | 58.8
|  | 0.0158 | 51.1
|  | 0.0119 | 44.2
|  | 0.0088 | 31.3
|  | 0.0063 | 25.1
|  | 0.0032 | 20.6
|  | 0.0014 | 6.6
Template Revision Date: |  | 2021-03-11 | [12 empty] | Page 1 of 1
Yes
No
Companion Measurements
Calibration Relationship
Note on F19: Pete Holter: This is used to format the temperature data below and to flag results documented beyond the resolution of the thermometer.
Note on B22: Pete Holter: Section 9.9: "Place the sedimentation specimen in the specimen mixing container and record the identification of the specimen mixing container"
Note on B23: Pete Holter: If there is enough material in the sedimentation sample, a separate sample is dried per D2216 to determine the water content. If not, the water content is determined in the container large enough to hold the soil suspension.
Note on B24: Pete Holter: If there is enough material in the sedimentation sample, a separate sample is dried per D2216 to determine the water content. If not, the water content is determined in the container large enough to hold the soil suspension.
Note on L24: Pete Holter: A different balance with a larger capacity may be needed for the original sample mass determination.
Note on O25: Pete Holter: The resolution is needed to format the original sample mass below and to flag results documented beyond the resolution of the balance.
Note on B37: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section 9.5: "Estimate the amount of moist mass needed for the sedimentation test..."
Note on B46: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section 9.8: "Record the mass of the moist soil, Mm, used for the sedimentation test..."
Note on J64: Pete Holter: Dry Mass of Specimen Plus Dispersant
Note on J65: Pete Holter: Dry Mass of Specimen
Note on J66: Pete Holter: Dry Mass of the Specimen Retained on the No. 200 Sieve
Note on J67: Pete Holter: Percent Passing the No. 200 Sieve

## Cylinder
Cylinder Checks
Cylinder ID: |  |  |  | Cyl-1 | Cyl-2
Name of Technician:
Date of Check
Procedure Used: |  |  |  | D7928-21 Annex A1 | D7928-21 Annex A1
Balance ID:
Thermometer ID:
Caliper ID:
Property | Tolerance
| Lower | Upper | Units | Data | Data
Height of Cylinder: |  |  | Approximately 457 mm | 431 | 447
Volume | Lower | Upper | Units | Data | Data
Temperature of Water To Be Added: | 15.0 / 59.00 | 30.9 / 87.62 | °C / °F | 25.3 | 17
Density of Water: |  |  | g/mL | 0.99697 | 0.99878
Fill to Calibration Mark
Mass of Water: |  |  | g | 997.3
Volume: | 995 | 1005 | mL | 1000.3 | Need Mass
Volume within Tolerance?: |  |  | Yes or No | Yes
Required Action: |  |  | Fill with Required Mass or None | None
Fill with Required Mass
Fill with This Much Water: |  |  | g | 996.97 | 998.78
Volume: | 995 | 1005 | mL | 1000 | 1000
Volume within Tolerance?: |  |  | Yes or No | Yes | Yes
Remarked at 1000 mL? |  |  | Yes or No | No | No
Vsp | 995 | 1005 | mL or cm3 | 1000 | 1000
Inside Area | Lower | Upper | Units | Data | Data
Measurement 1 - Distance to 1000 mL Mark: |  |  | mm | 345.4 | 360
Measurement 1 - Distance to 1000 mL Mark: |  |  | mm | 345.2 | 360
Measurement 1 - Distance to 1000 mL Mark: |  |  | mm | 345.7 | 360
Average - Distance to 1000 mL Mark: |  |  | mm | 345 | 360
Ac |  |  | cm2 | 29 | 27.8
Calculated Inside Diameter: | 55 | 64 | mm | 61 | 59
Area Uniformity | Lower | Upper | Units | Data | Data
Mass of Water: | 149 | 151 | g | 148 | 148
Temperature of Water: | 15.0 / 59.00 | 30.9 / 87.62 | °C / °F | 25 | 60
Density of Water: |  |  | g/mL | 0.99705 | 0.99901
Area Uniformity Lower Limit: |  |  | mm | 48.6 | 50.6
Area Uniformity: |  |  | mm | 49 | 56
Area Uniformity Upper Limit: |  |  | mm | 53.7 | 56
Area Uniformity Pass or Fail: |  |  | Pass or Fail | Pass | Pass
Template Revision Date: | 2021-02-15 |  |  |  | Page 1 of 1
Note on B7: Pete Holter: It is assumed that the balance is readable to 0.01 g.
Note on B8: Pete Holter: It is assumed that the thermometer is readable to 0.1 degree.
Note on B9: Pete Holter: It is assumed that the caliper is readable to 0.01 mm.
Note on B16: Pete Holter: Previous versions of D7928 did not have you do it this way.
Note on B21: Pete Holter: This is what previous versions of D7928 required you to do. The new version just tells you to remark the cylinder if the fill to calibration mark fails. This would be an acceptable method in that case.
Note on E22: Pete Holter: Calculated from Temperature
Note on E23: Pete Holter: The volume is within tolerance if the water level is within about 2 mm of the calibration mark.

## Hydrometer
Hydrometer Checks
Hydrometer ID: |  | Hydro-1 | Hydro-2 | Hydro-3
Hydrometer Type: |  | 152H | 152H | 152H
Initial Check
Name of Technician:
Date of Check
Procedure Used: |  | D7928-21 Annex A1 | D7928-21 Annex A1
Balance ID:
Thermometer ID:
Caliper ID:
Length Dimensions | Tolerance and Unit | Data | Data | Data
HL: | 275- 285 mm | 275 | 285
Hb: | 243 - 247 mm | 241 | 248
HLb: | 115 - 142 mm | 135.9 | 134
Dm: | 29 - 32 mm | 28 | 33
Hcb: | 58 - 71 mm | 67.3 | 70
Hs: | 81 - 85 mm | 82 | 82
Stem Length Above Scale: | > 15 mm | 16 | 16
Stem Length Below Scale: | > 3 mm | 5 | 4
r1 | 0.995 for 151H, -5 for 152H | -5 | -5
r2 | 1.038 for 151H, 60 for 152H | 60 | 60
Scale Key Reference Point | Tolerance and Unit | Data | Data
Reference Point Hydrometer Reading: |  | 0 | -1
Reference Point Temperature: | 15.0 - 30.9 °C | 23 | 25
Scale Key Reference Point: | 151H: 0.999 - 1.001, 152H: -1 - 1 g/L | 1.08 | 0.86
Temperature-Density Correction | Calibration Relationship | Data | Data
Trial 1 - Hydrometer Reading | dimensionless for 151H and g/L for 152H | 7 | 1.0055
Trial 1 - Temperature | °C | 17.8 | 18.4
Trial 2 - Hydrometer Reading | dimensionless for 151H and g/L for 152H | 6.5 | 1.0052
Trial 2 - Temperature | °C | 19.1 | 20.3
Trial 3 - Hydrometer Reading | dimensionless for 151H and g/L for 152H | 6 | 1.005
Trial 3 - Temperature | °C | 20.6 | 22.8
Trial 4 - Hydrometer Reading | dimensionless for 151H and g/L for 152H | 5.5 | 1.0048
Trial 4 - Temperature | °C | 21.9 | 23.7
Trial 5 - Hydrometer Reading | dimensionless for 151H and g/L for 152H | 5 | 1.0032
Trial 5 - Temperature | °C | 23.3 | 27.3
Standard Deviation | dimensionless for 151H and g/L for 152H | 0.1 | 1.3
A/B | 151H: A, 152H: B | 9.6 | 5.4
Meniscus Correction |  | Data | Data
Top of Meniscus: |  | 3 | 0.9995
Bottom of Meniscus: |  | 7 | 1
Cm |  | -4 | -0.0005
Volume of the Hydrometer Bulb | Tolerance and Unit | Data
Method of Determination: | Direct or Indirect | 1 | 1
Trial 1 - Initial Reading: | Trials within 2 mL or g | 0 | 150
Trial 1 - Final Reading: | Trials within 2 mL or g | 56 | 204
Trial 2 - Initial Reading: | Trials within 2 mL | 2 | 150
Trial 2 - Final Reading: | Trials within 2 mL | 60 | 206
Trial 3 - Initial Reading: | Trials within 2 mL | 0 | 150
Trial 3 - Final Reading: | Trials within 2 mL | 56 | 204
Vhb | 50 - 72 mL/cm3 | 57 | 55
1/2 Vhb for cb: |  | 28.5 | 27.5
hr1 | mm | 174 | 186
hr2 | mm | 67.5 | 70
Page 1 of 2
12 Month Checks
Name of Technician:
Date of Check
Procedure Used: |  | D7928-21 Annex A1 | D7928-21 Annex A1
Caliper ID:
Scale Key Reference Point | Tolerance and Unit | Data | Data
Reference Point Hydrometer Reading:
Reference Point Temperature: | 15.0 - 30.9 °C
Scale Key Reference Point: | 151H: 0.999 - 1.001, 152H: -1 - 1 g/L
Name of Technician:
Date of Check
Procedure Used: |  | D7928-21 Annex A1 | D7928-21 Annex A1
Caliper ID:
Scale Key Reference Point | Tolerance and Unit | Data | Data
Reference Point Hydrometer Reading:
Reference Point Temperature: | 15.0 - 30.9 °C
Scale Key Reference Point: | 151H: 0.999 - 1.001, 152H: -1 - 1 g/L
Template Revision Date: | 2025-08-06 |  |  | Page 2 of 2
Direct
Indirect
Direct
Indirect
Note on B9: Pete Holter: It is assumed that the balance is readable to 0.01 g.
Note on B10: Pete Holter: It is assumed that the thermometer is readable to 0.1 degree.
Note on B11: Pete Holter: It is assumed that the caliper is readable to 0.01 mm.
Note on B12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section 10.1: "Check and record the dimensions of the 151H or 152H hydrometers as presented in Annex A1..." However, Hr1 and Hr2 are included below with the bulb volume determination.
Note on B13: Pete Holter: Total length of the hydrometer
Note on B14: Pete Holter: The distance from the key reference point on the nominal scale to the bottom of the hydrometer bulb
Note on B15: Pete Holter: The overall length of the hydrometer bulb
Note on B16: Pete Holter: Maximum diameter of the bulb
Note on B17: Pete Holter: The distance from the bottom of the hydrometer bulb to the maximum diameter of the bulb
Note on B18: Pete Holter: The length of the quoted nominal scale as stated in Specification E100
Note on B19: Pete Holter: The length of stem above the topmost graduation
Note on B20: Pete Holter: The length of stem below the lowest graduation
Note on B21: Pete Holter: Minimum hydrometer reading
Note on B22: Pete Holter: Maximum hydrometer reading
Note on B23: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "Points—The 152H hydrometer should read 0.0 +/- 1 g/L and 1.000 +/- 0.001 specific gravity for the 151H hydrometer at the bottom of the meniscus when placed in distilled water that is free of gas bubbles at the calibration temperature of the hydrometers."
Note on B25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "At temperatures other than the calibration temperature of the hydrometer, the required reading shall be adjusted..."
Note on C25: Pete Holter: Temperatures above 40 are treated as Fahrenheit.
Note on B27: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section 10.2.2: "This option generates a general calibration relationship between the hydrometer reading of a control cylinder filled with the reference solution and the temperature, which eliminates the need for companion measurements during the test."
Note on B38: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sections 10.2.2.1 and 10.2.2.2: "The standard deviation of the 5 [hydrometer] measurements shall be less than 0.0005 [for a 151H hydrometer] and less than 0.5 g/L [for a 152H hydrometer.]"
Note on B39: Pete Holter: The average value used when computing the temperature-density correction
Note on B40: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section 10.3: "Hydrometers are designed to be read at the fluid surface..."
Note on B45: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The volume determination required by Sections 10.4.1 and A1.3.4 does not necessitate three trials for the indirect method. However, Figure X1.4 shows the volume being determined by the indirect method, and includes data for three trials. In case the intention of Section A1.3.4 is to require three trials for either method, this spreadsheet requires three trials for both methods.
Note on B47: Pete Holter: This row receives an initial reading in the graduated cylinder for the direct method, or 0 as a tare for the indirect method.
Note on B48: Pete Holter: This row receives a final reading in the graduated cylinder for the direct method, or a mass of displaced water for the indirect method.
Note on B49: Pete Holter: This row receives an initial reading in the graduated cylinder for the direct method.
Note on B50: Pete Holter: This row receives a final reading in the graduated cylinder for the direct method.
Note on B51: Pete Holter: This row receives an initial reading in the graduated cylinder for the direct method.
Note on B52: Pete Holter: This row receives a final reading in the graduated cylinder for the direct method.
Note on B53: Pete Holter: The volume of the hydrometer bulb
Note on B54: Pete Holter: Provided as a guide to help determine where to mark the center of (volume) buoyancy (cb)
Note on B55: Pete Holter: Distance from the minimum hydrometer reading to the center of (volume) buoyancy (from r1 to cb)
Note on B56: Pete Holter: Distance from the maximum hydrometer reading to the center of (volume) buoyancy (from r2 to cb)
Note on B65: Pete Holter: It is assumed that the caliper is readable to 0.01 mm.
Note on B66: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "Points—The 152H hydrometer should read 0.0 +/- 1 g/L and 1.000 +/- 0.001 specific gravity for the 151H hydrometer at the bottom of the meniscus when placed in distilled water that is free of gas bubbles at the calibration temperature of the hydrometers."
Note on B68: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "At temperatures other than the calibration temperature of the hydrometer, the required reading shall be adjusted..."
Note on C68: Pete Holter: Temperatures above 40 are treated as Fahrenheit.
Note on B73: Pete Holter: It is assumed that the caliper is readable to 0.01 mm.
Note on B74: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "Points—The 152H hydrometer should read 0.0 +/- 1 g/L and 1.000 +/- 0.001 specific gravity for the 151H hydrometer at the bottom of the meniscus when placed in distilled water that is free of gas bubbles at the calibration temperature of the hydrometers."
Note on B76: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Section A1.3.3: "At temperatures other than the calibration temperature of the hydrometer, the required reading shall be adjusted..."
Note on C76: Pete Holter: Temperatures above 40 are treated as Fahrenheit.

## Table 1
Table 1 Mimimum Dry Mass Requirements
Maximum Particle Size of Material (99% or more passing) |  | Minimum Dry Mass
Sieve Designation | Particle Size, mm
6 in. | 152.4 | 500000 | g
3 in. | 76.2 | 70000 | g
1 in. | 25.4 | 3000 | g
3⁄4 in. | 19.1 | 1300 | g
3⁄8 in. | 9.5 | 165 | g
No. 10 | 2 | 50 | g
Template Revision Date: | 2020-03-02 | Page 1 of 1

## Water Density
Table A1.1
Temperature (T) Versus Density of Water (ρw)
Temperature in °C | Water Density (ρw) in g/ml
15 | 0.9991
15.1 | 0.99909
15.2 | 0.99907
15.3 | 0.99906
15.4 | 0.99904
15.5 | 0.99902
15.6 | 0.99901
15.7 | 0.99899
15.8 | 0.99898
15.9 | 0.99896
16 | 0.99895
16.1 | 0.99893
16.2 | 0.99891
16.3 | 0.9989
16.4 | 0.99888
16.5 | 0.99886
16.6 | 0.99885
16.7 | 0.99883
16.8 | 0.99881
16.9 | 0.99879
17 | 0.99878
17.1 | 0.99876
17.2 | 0.99874
17.3 | 0.99872
17.4 | 0.99871
17.5 | 0.99869
17.6 | 0.99867
17.7 | 0.99865
17.8 | 0.99863
17.9 | 0.99862
18 | 0.9986
18.1 | 0.99858
18.2 | 0.99856
18.3 | 0.99854
18.4 | 0.99852
18.5 | 0.9985
18.6 | 0.99848
18.7 | 0.99847
18.8 | 0.99845
18.9 | 0.99843
19 | 0.99841
19.1 | 0.99839
19.2 | 0.99837
19.3 | 0.99835
19.4 | 0.99833
19.5 | 0.99831
19.6 | 0.99829
19.7 | 0.99827
19.8 | 0.99825
19.9 | 0.99823
20 | 0.99821
20.1 | 0.99819
20.2 | 0.99816
20.3 | 0.99814
20.4 | 0.99812
20.5 | 0.9981
20.6 | 0.99808
20.7 | 0.99806
20.8 | 0.99804
20.9 | 0.99802
21 | 0.99799
21.1 | 0.99797
21.2 | 0.99795
21.3 | 0.99793
21.4 | 0.99791
21.5 | 0.99789
21.6 | 0.99786
21.7 | 0.99784
21.8 | 0.99782
21.9 | 0.9978
22 | 0.99777
22.1 | 0.99775
22.2 | 0.99773
22.3 | 0.9977
22.4 | 0.99768
22.5 | 0.99766
22.6 | 0.99764
22.7 | 0.99761
22.8 | 0.99759
22.9 | 0.99756
23 | 0.99754
23.1 | 0.99752
23.2 | 0.99749
23.3 | 0.99747
23.4 | 0.99745
23.5 | 0.99742
23.6 | 0.9974
23.7 | 0.99737
23.8 | 0.99735
23.9 | 0.99732
24 | 0.9973
24.1 | 0.99727
24.2 | 0.99725
24.3 | 0.99723
24.4 | 0.9972
24.5 | 0.99717
24.6 | 0.99715
24.7 | 0.99712
24.8 | 0.9971
24.9 | 0.99707
25 | 0.99705
25.1 | 0.99702
25.2 | 0.997
25.3 | 0.99697
25.4 | 0.99694
25.5 | 0.99692
25.6 | 0.99689
25.7 | 0.99687
25.8 | 0.99684
25.9 | 0.99681
26 | 0.99679
26.1 | 0.99676
26.2 | 0.99673
26.3 | 0.99671
26.4 | 0.99668
26.5 | 0.99665
26.6 | 0.99663
26.7 | 0.9966
26.8 | 0.99657
26.9 | 0.99654
27 | 0.99652
27.1 | 0.99649
27.2 | 0.99646
27.3 | 0.99643
27.4 | 0.99641
27.5 | 0.99638
27.6 | 0.99635
27.7 | 0.99632
27.8 | 0.99629
27.9 | 0.99627
28 | 0.99624
28.1 | 0.99621
28.2 | 0.99618
28.3 | 0.99615
28.4 | 0.99612
28.5 | 0.99609
28.6 | 0.99607
28.7 | 0.99604
28.8 | 0.99601
28.9 | 0.99598
29 | 0.99595
29.1 | 0.99592
29.2 | 0.99589
29.3 | 0.99586
29.4 | 0.99583
29.5 | 0.9958
29.6 | 0.99577
29.7 | 0.99574
29.8 | 0.99571
29.9 | 0.99568
30 | 0.99565
30.1 | 0.99562
30.2 | 0.99559
30.3 | 0.99556
30.4 | 0.99553
30.5 | 0.9955
30.6 | 0.99547
30.7 | 0.99544
30.8 | 0.99541
30.9 | 0.99538
Template Revision Date: | 2020-03-02
Page 1 of 1