## 6-in. Proctor Mold Check
6-inch Diameter Proctor Mold Volume Determination per D698-12/D1557-12 Annex
| Mold ID: |  |  |  | Name Technician:
| Caliper ID: |  |  |  | Date of Check:
| Thermometer ID: |  |  |  | Procedure: |  | D698-12/D1557-12 Annex
Balance ID:
Linear Measurement Volume Determination | [5 empty] | Water-Filling Volume Determination
Set the Balance To Read in Grams
[7 empty] | and Enter the Resolution Here: | 0.01
|  | Set the Calipers To Read in Inches and Enter the Resolution Here: | 0.0005 |  |  |  | Enter the Resolution of the Thermometer Here: | 0.1
Measurements | Internal Diameter (Top) | Internal Diameter (Bottom) | Measurements | Height |  |  |  | Trial 1 | Trial 2
M1 | 6.0065 | 5.997 | M1 | 4.554 |  | Mass of Mold, Base Plate, Grease, Glass Plate (g): |  | 7205.31
M2 | 5.997 | 5.9975 | M2 | 4.554 |  | Mass Filled with Water (g): |  | 9321.16
M3 | 5.9885 | 5.9975 | M3 | 4.554 |  | Mass of Water (g): |  | 2116
M4 | 5.994 | 5.985 | M4 (optional) |  |  | Temperature of Water (<= 40 = C; > 40 = F): |  | 23
M5 | 5.996 | 5.991 | M5 (optional) |  |  | ASTM Density of Water (g/cm3): |  | 0.99754
M6 | 6.0045 | 5.995 | M6 (optional) |  |  | AASHTO Density of Water (lb/ft3): |  | 62.274
[6 empty] | ASTM Volume (cm3): |  | 2121
[6 empty] | ASTM Volume (ft3): |  | 0.0749
Average Diameter: | 5.996 |  | Average Height: | 4.554 |  | AASHTO Volume (ft3): |  | 0.0749
ASTM Average Water-Filling Volume (ft3):
Linear Measurement Volume (ft3): |  | 0.0744 |  |  |  | AASHTO Average Water-Filling Volume (ft3):
Is the Average Diameter within 6.000 +/- 0.026 in? |  |  | Yes |  |  | ASTM: Is the Volume Determined By the Water-Filling
[6 empty] | Method 0.0750 +/- 0.0009 ft3? |  |  | Yes
Is the Average Height within 4.584 +/- 0.018 in? |  |  | No
AASHTO: Is the Volume Determined By the Water-Filling
Is the Volume Determined By the Linear Measurement Method 0.0750 +/- 0.0009 ft3? |  |  | Yes |  |  | Method 0.0750 +/- 0.0009 ft3? |  |  | Yes
ASTM Linear Measurement to Water-Filling Volume Comparison
|  | Volume Comparison in %: |  |  |  | 0.67
|  | Within 0.5% nominal? |  |  |  | No
Standardized Volume for Calculating Maximum Density
|  | ASTM Linear Measurement: |  |  |  | 0.0744
|  | ASTM Water-Filling: |  |  |  | Requires 2 Trials
|  | ASTM Average: |  |  |  | 0.0747
|  | AASHTO Water-Filling: |  |  |  | 0.0749
Note: Nonconformities are highlighted in bold red.
Preparation/Revision Date: 6/8/18
Note on I6: Pete Holter: Also uses the water density table in T 19 for conformance with AASHTO T 99/T 180.

## Water Density Table
TABLE 3 Density of Water
Temperature |  | kg/m3 | lb/ft3
C | F
15.6 | 60 | 999.01 | 62.366
18.3 | 65 | 998.54 | 62.336
21.1 | 70 | 997.97 | 62.301
23 | 73.4 | 997.54 | 62.274
23.9 | 75 | 997.32 | 62.261
26.7 | 80 | 996.59 | 62.216
29.4 | 85 | 995.83 | 62.166